# GW Instek GSP-9300頻譜分析儀 FSK解調及分析功能DEMO

這次的演示是GSP-9300 頻譜分析儀的FSK解調及分析功能。

這 DEMO KIT是個ASK及FSK的信號產生器。此DEMO KIT有3種預設頻率，3種波特率及可設定輸出為連續或叢發模式。

架設方式：

- 連接DEMO KIT到GSP-9300前面板的USB接口

- 用SMA RF CABLE連接DEMO KIT輸出端到GSP-9300的輸入端。

儀器操作:

- 首先設定DEMO KIT的參數.

- 按Option Control鍵 > Demo Kit[F4]鍵>Demo Kit[F1]鍵*2次,啟動DEMO KIT.

- 按Modulation [F2]鍵, 選擇FSK[F2]鍵

- 按Frequency[F3]鍵,選擇433MHz[F2]鍵

- 按Deviation[F4]鍵,選擇50kHz[F2]鍵

- 按Baudrate[F5]鍵, 選擇波特率為4800[F2]鍵

- 按Output Mode[F6]鍵, 選擇10101010字串,Seq5[F7].

- DEMO KIT的參數設定結束.

- 按 Frequency 鍵> Center Freq[F1]鍵>4,3,3,MHz[F2]鍵, 設定中心頻率為433MHz

- 按 Span鍵> Span[F1]鍵 >5,0,0,kHz[F3]鍵, 設定頻段為500kHz

- 在頻譜分析儀上可以觀測到FSK信號.

- 開始設定FSK解調及分析功能觀測FSK的信號.

- 按 Measure鍵> Demod[F2]鍵>FSK[F5]鍵>FSK Analysis[F1]*2次,啟動FSK分析功能

- FSK分析測量參數，包括頻偏，調製速率，載波功率，載波頻率偏移，信納比，符號碼(symbol)及波形(waveform)。用戶可以設定頻偏，載波功率和載波偏移的合格/失敗測試結果。

- 按Symbol Rate[F2]鍵，輸入4,8,0,0,s [F2]，設定符號碼率為4800

- 按Depth Axis[F6]鍵，按Scale/Div[F3]鍵，輸入2，0，kHz[F3]. 設定每格為20kHz.

- 按Return[F7]

- 目前螢幕上半部的測試結果為FSK解調的波形(waveform).接著可以按View[F4]鍵，選擇符號碼(symbol)[F1].

- 用戶可以在螢幕上看到FSK解調的符號碼為之前設定的Seq5,10101010字串.

- 按Return[F7]鍵 > 按More　1/2[F7]鍵 > 按Limit[F4]鍵

- 按Freq.Deviation[F2],設定5,0,kHz[F3]

- 按Carrier offset[F3],設定１,kHz[F3]

- 按Limit[F1]*2次，啟動合格/失敗結果測試 > 按Return[F7]*2次

- 按View[F4]鍵，選擇波形(waveform)[F2] >按Return[F7] > 按More 1/2[F7]

- 按AF Trigger[F1]>Trigger mode[F3]*2次,設定單次觸發>Trigger Level[F4],設定觸發準位為1,0,kHz[F3]

- 按Action Now[F6], 啟動單次觸發。GSP-9300暫停測量FSK信號。

- 以上是GSP-9300 FSK分析功能的介紹.

- 感謝您的收看

# GW Instek GSP-9300 Spectrum Analyzer FSK Demodulation and Analysis Function DEMO

Now, I am going to demonstrate the FSK demodulation and analysis function of GSP-9300 spectrum analyzer.

This demo kit is an ASK and FSK signal generator. In this demo kit, there are three preset frequency, three Baud rates and selectable continuous or burst output mode.

First of all, connect demo kit with the USB port located at the front panel of GSP-9300. Next, connect demo kit output terminal with GSP-9300 input terminal with SMA RF cable.

- To begin the operation, first, input demo kit parameters.

- Press option control key, demo kit (F4) key, demo kit (F1) key twice to activate demo kit

- Press modulation (F2) key and select FSK (F2) key

- Press frequency (F3) key and select 433 MHz (F2) key

- Press deviation (F4) key and select frequency deviation of 50kHz (F2) key

- Press Baud rate (F5) key and select baud rate of 4800 (F1) key

- Press output mode (F6) key and select binary code 10101010,Seq5(F7)

- This is the end of demo kit parameters input. Next, set up the FSK demodulation & analysis function to observe FSK signals.

- Press frequency key, center freq (F1) key, input 4, 3, 3, MHz (F2) key to set center frequency as 433 MHz.

- Press span key, span (F1) key, input 5, 0, 0, kHz (F3) key to set frequency bandwidth as 500kHz.

- Now, FSK signals are shown on the screen.

- Next, set FSK demodulation and analysis function to observe FSK signals.

- Press measure key, Demod (F2) key, FSK Analysis (F5) , FSK Analysis (F1)twice to activate FSK analysis function.

- FSK Signal Analysis measures parameters including frequency deviation, modulation rate, carrier power, carrier frequency offset, Signal-to-noise and distortion ratio (SINAD), symbol and waveform. Users can set the criterion in frequency deviation, carrier power and carrier offset for pass/fail test result determination.

- Press symbol rate(F2) key, input 4,8,0,0 s (F2) to set symbol rate as 4800

- Press depth Axis (F6) key, and Scale/Div(F3), input 2,0 kHz(F3) to set 20kHz per div

- Press return (F7) key

- Now, the top half of the screen shows the test result of waveform demodulated by FSK. Next press View (F4) and select symbol (F1)

- Now, the previous set FSK demodulation symbol of binary code Seq5,10101010 is shown on the screen

- Press Return(F7), and More (F7), press Limit(F4)

- Press Freq. Deviation(F2),input 5,0,kHz(F3)

- Press Carrier offset(F3),input 1,kHz(F3)

- Press Limit(F1) twice to activate pass/fail test result, press Return(F7) twice

- Press View(F4), select waveform(F2), press Return(F7), and More (F7)

- Press AF Trigger (F1), Trigger mode (F3) twice, Trigger level (F4), input 1,0 Enter (F3) to set AF Trigger and Trigger Level as 10kHz

- Press Action Now (F6) to activate AF Trigger. GSP-9300 stops measuring FSK signals.

- This concludes the demonstration of GSP-9300 FSK analysis function.

- Thank you very much for your interest in GSP-9300.